## Slide 1
The 224 Liberas ( 32 cells x 7 stations ) are now all doing Closed Orbit
measurement and slow Orbit Correction
the T-b-T capability has been tested and
is close to be fully commissioned
the FA (10KHz) distribution is also in progress
of commissioning
7 cells had been done BEFORE Winter shut-down (in 3 MDTs)
13 cells were done DURING Winter shut-down
12 more cells were done AFTER Winter shut-down (in 4 MDTs)
the new Electron Beam Position Measurement System (BPM)

## Slide 2
10
um
20 Cells with LIBERA-BPMs
4 Cells with
Old-BPMs
8 Cells with
Old-BPMs
Beam Orbit Plot (i.e. substraction from a reference orbit)
Also : The strategy of dealing with new / old offsets appears globally
successful in avoiding shifts of the X-ray beam for the beamlines
an intermediate view . . . . .

## Slide 3
a BPM system is not only the Libera . . .
Libera
Control
Feedback
or Orbit-
Stabilizer
firmware
release
224
device
Servers
fast links
Brilliance
1.82 but soon 2.02 . . . . .
and then to 2.04.1 . . . . .
and committed to 2.2 . . .
Tango
still under modified
‘Soleil’ server, but
soon to be replaced
with a completely
restructured server
C.C.
integrated
by I-tech
mainly
‘nasty’
fillings
protection for
peak voltages
only 1.82 features
implemented, the
full 2.04.1 features
to be added after
Interlock &
Post-Mortem
more flexible
trigger system
trouble-shooting poor
RF interconnections,
compensating for unequal
amplitudes & phases
100Mbit
Ethernet
+

## Slide 4
a BPM system is 3 worlds meeting together . . .
Libera
Feedback
or Orbit-
Stabilizer
firmware
release
device
Servers
Interface with
Control system
+
RF signals
1
2
3

## Slide 5
Failure rate of the Liberas sofar :
We did very simple Site-Acceptance-Tests, ‘in-final-situ’
i.e. no precise Lab measurements to verify the ultimate specs
Reasons : 1) confidence in I-Tech’s FATs
2) reduction of work-load
3) concentrate efforts on finding the real crazy units
Out of 230 delivered units :
a total of 15 (3 shipments of 5, 7 and 3 units) needed repair at I-Tech’s site
On top of these 15 we also had some ( ~10 ) software crashes inside the Libera
that could be repaired (on-site) in 3 different ways :
1) I-Tech logged-on to the unit, diagnosed, and solved the problem.
2) connect a Laptop, run some tools and get it alive again.
take an image of the flash memory of a good unit (with a Laptop),
download this into a totally blocked Libera to get it alive again.
Sometimes such blockages happen during a firmware upgrade.

## Slide 6
1m thermo-resistant
RF cables
BPM
Block
survey monument
bad Yoke . . .

## Slide 7
BPM
Block
1m thermo-
resistant
RF cables
interconnection to
20 m long RF cables

## Slide 9
Too HOT !
WARM
HOT !

## Slide 10
220 V
SR_clock
per Libera :
20 RF interconnections
timing interconnections
Ethernet interconnections
in total : ~7000 interconnections for the whole upgrade . . . .
Post
Mortem
Lib.
Int.
RF
atten.
&
BP filter
Interlock
Ethernet
Inject_Trig.
Event_Trig.
Libera
1m thermo-resistant
RF cables
20m RF transmission cables
short RF
transmission
cables
-3dB 40MHz

## Slide 11
-3dB
40MHz
~25m RG223
~7dB att.
60cm RG316
5 Euros
F
M
M
F
1m
RF signal chain : from pick-up button to ADC
128 Euros
ADC
strong offsets
beam - BPMblock
thus :
unequal Ampl.
+9 % [average]
+45 % [worst case]
different cable lenghts
thus :
unequal Phases
21 mm [average]
106 mm [worst case]
low-pass
filter
band-pass
SAW filter

## Slide 12
DSC means Digital Signal Conditioning,
its implies the multiplexing & de-mux. of the 4 signals
through the 4 RF & ADC chains to re-calibrate the Libera
It is done at exactly fsr / 27 (~13KHz)
This mechanism is in principle transparent and invisible
in the FA (10KHz) and SA (10Hz) outputs

## Slide 13
Turn-by-Turn data : 22KHz component of ~15um pk-p in Vert. Plane,
+ spikes due to DSC Libera-switching (at 27 TbT-intervals)

## Slide 14
take TbT data  FFT  assess spike amplitude
then compare for each station :
Spike-amplitude versus Un-balance in signal-amplitude
&
Spike-amplitude versus Un-balance in signal-phase
Conclusion : clear relation & proportionality between the
undesired spikes and the ‘in-perfections’ in RF signals
So the Libera has some problems with ‘real’ RF signals,
However, it offers solutions :
Offset-tune
Spike removal feature (2.00)
We tested these on 8 prototype units in early 2008 :
looking at FA data (and its spectrum)
under ‘nasty’ filling patterns

## Slide 15
Integrated noise spectrum
Hz
nm

## Slide 16
Conclusion :
a large offset-tune
will move all
spurious lines
out of the
bandwidth of
interest of
the FA output

## Slide 17
-3dB
40MHz
~25m RG223
~7dB att.
60cm RG316
5 Euros
F
M
M
F
1m
RF signal chain : from pick-up button to ADC
128 Euros
ADC
strong offsets
beam - BPMblock
thus :
unequal Ampl.
+9 % [average]
+45 % [worst case]
different cable lenghts
thus :
unequal Phases
21 mm [average]
106 mm [worst case]
low-pass
filter
band-pass
SAW filter

## Slide 18
Drift in 4 x 10mA : with & without the filters in front of Libera
10um

## Slide 19
Drift in 4 x 10mA : with & without the filters in front of Libera

## Slide 20
The drift of a Libera with signal strength :
specified to <1um, but true for ‘nice’ input signal (phase, amplitude)
not for strongly un-balanced input signals

## Slide 21
the modified DDC filter works
well i.e. the MAF design
Situation with standard DDC filter
‘smearing’
but only one design can be installed at a time . . .
Improving the installation procedure (i.e. reducing time) would help a lot . . .

## Slide 22
Time-Scanning of the Filter synchronization with the 1/3 fill
BEAM (33%)
Filter

## Slide 23
Turns
Sum
Signal
(TbT)
of all
Liberas
Sum
Signal
(TbT)
of 2
Liberas
strong losses . . .
Injection (from 0)

## Slide 24
Sum
Signal
(TbT)
of 2
Liberas
with standard
DDC filter
i.e. with some
smearing
differentiated
Sum Signals

## Slide 26
Sum
Signal
(TbT)
of all
Liberas
(mean)
Sum
Signal
(TbT)
of all
Liberas
Turns
224 BPMs

## Slide 27
Sum
Signal
(TbT)
of all
Liberas
(mean)
X position
(TbT, mm)
of a Libera

## Slide 28
X position
(TbT, mm)
of a Libera
X position
(TbT, mm)
of ALL
Liberas
224 BPMs
Turns

## Slide 29
The reliability of a BPM value : survey the incoherency (or Q-value)
a station is discarded from Orbit control if the incoh drifts too much,
but now we have to sharpen our maintenance tools to track the cause
A
B
C
D
incoherency

## Slide 30
from the SA-Sum
signal of each
Libera a Life-time
measurement
can be done :
looks very
promising !!

## Slide 31
looks very
promising !!
Also : it is
a different and
complementary
way of checking
your BPM
signals :
anomalous
stations will
rapidly show-up

## Slide 32
additional degradation with partial & single bunch fills :
when should the Libera do its DSC switching ??
output after the 40MHz wide BPF = typical Libera RF input @single bunch fill
here ?
or here ?

## Slide 33
Another Libera artefact appears in SA data with offset-tuning :
a slow 200nm pk-pk oscillation, the period is determined by the value of the offset
data taken under 16 bunch, ‘perfect stable’ beam (Sum & Split)
offset=10units
offset=20units

## Slide 34
The slow 200nm pk-pk oscillation in SA data can be minimized (suppressed)
by adjusting the switching delay parameter,
data taken under 16 bunch, ‘perfect stable’ beam (Sum & Split)
‘bad’ switching delay
‘good’ switching delay

## Slide 35
scanning the switching delay value & measuring the spike amplitudes at each step :
the individual good / bad values can be easily determined for each Libera
the 224 BPMs
20
steps
of 4
units

## Slide 36
Planning of remaining tasks :
Firm- and software wise : - finish installation of 2.02 release
- install the 2.04.1 release
- fully test & commission new device-server
- fully test & commission the ‘all’ server
- add functionalities available under 2.04.1
BPM functionality wise : - re-commission Interlock (incl. Post-Mortem)
- fully commission Turn-by-Turn mode :
- with standard DDC filter
- with MAF filters
Maintenance wise : - store & retrieve suitable statistics in/from
data-base in efficient way to detect &
identify quickly problems (drifts, fluct.)
- define functionalities for 2.00 release
Fast-Feedback wise : - get all FA data-streams connected,
collected and analyzed, in preparation of
future new Orbit Stabilization system

## Slide 37
The ESRF BPM system has perhaps 2 half-sisters . . . :
Feedback
or Orbit-
Stabilizer
device
Servers
ALBA
+
PETRA-3
Libera
Brilliance
RF signals imposed by
circumference & fillings
Tango control system

## Slide 38
The ESRF wants to keep close communication
with all members of the (diverging. . .) Libera community,
i.e. with its half-sisters . . .
but equally with distant nephews . . .
and hopefully new . . . children !
So let’s stay in touch !