## Slide 1: 3 years of experience of Libera BPMs at SOLEIL
Nicolas HUBERT
Synchrotron SOLEIL
On behalf of Diagnostics group

## Slide 2: Libera Status
161 units of Libera Electron
  Delivered in 2005
  In operation for 3 years
  Distribution:
    Storage ring => 120 +1 units
    Booster => 22 units
    TL2 => 3 units
    TL1 => 1 unit
    Spares and tests => 14 units
Release 1.82 installed with modified FPGA design for FOFB application

## Slide 3: Configuration for User Operation
Activated Data sources:
  SA => To monitor and correct the orbit (slow orbit feedback)
  FA => To correct the orbit (fast orbit feedback) and monitor noise spectrum
  TT (standard filter)
Interlock
  Vertical plane:
    46 BPMs involved
    Thresholds from ±0.3 mm (IR mirrors) up to ± 1mm (bending magnets)
  Horizontal plane:
    8 BPMs involved
    Threshold at ±0.8 mm (IDs)
  Gain dependant (Inhibited if gain < -40 dBm)
Automatic Gain Control always activated
Post-Mortem on all BPMs (16 ksamples)
  Triggered by the Machine Interlock Signal
Detuning of ~9 kHz of sampling frequency (offset tune = 220)
External source for switching

## Slide 4: SOLEIL Releases Policy
Release 1.82 should fit SOLEIL Requirements:
  Difficulties with MAF have to be understood and solved
  Beam Current Dependence Tables (Libera Electron) have to be tested and filled.
Large amount of work for each upgrade
  Upgrade process can be very long if Linux system is updated ( 1.46 to 1.60 and 1.82 to 2.0)
  Integration and testing of FOFB application in the new FPGA design
  Testing of all functionalities ?

## Slide 5: SOLEIL Releases Policy
SOLEIL upgrade to 1.82:
After 2 days of operation, systematic beam losses when trying to increase the current of stored beam:
  For each new release, improvements have to be compared to the upgrade time, testing time and remaining bug risk.
  At the moment SOLEIL has decided to stay with release 1.82.
RESTART
Systematic beam losses when stored beam reaches ~50 mA
  Level of Sum value had changed from release 1.62 to 1.82
  Threshold for DSC coefficients computation not modified accordingly
    ► Wrong DSC coefficients applied creating orbit jumps and interlocks.
    ► Several hours of beam time have been lost

## Slide 6: Moving Average Filter with Libera Electron
New Turn By Turn Filter designed by K. Scheidt (ESRF) and implemented in Libera package from release 1.82.
No more smearing between turns:
Standard Filter
MAF Filter

## Slide 7: Moving Average Filter with Libera Electron
Time-Scanning of the Filter synchronization:
BEAM (33%)
Filter
Brilliance
S/N=9000
Electron
S/N=4
Signal over noise ratio is very low for Libera Electron:
  Automatic routine to scan and find the appropriate MAF delay is not possible
    MAF Delay is defined roughly with automatic routine and then adjusted manually for each BPM….Time consuming!
Courtesy K. Scheidt

## Slide 8: Moving Average Filter with Libera Electron
Turn by Turn acquisition after MAF Length optimization
1 Turn detection
Standard filter
MAF filter

## Slide 9: Hardware Failures History
20 Libera have presented failure:
Maintenance
  Libera are turned ON 365/24 since 2006
    Some Libera are becoming noisy (fans wear)
      -> Do we have an idea of their lifetime?
      -> Have to foresee a fan replacement campaign at the first signs of failure.