## Slide 1
Annual Review and Performance Improvements of DLS RF System
Pengda Gu
on behalf of
Diamond Storage Ring RF Group
15th ESLS-RF Meeting, October 5-6th, ESRF

## Slide 2
Agenda
Operational statistics
Fault analysis and treatment
Pulse blanker test
Cavity 2 repair progress
Multipactor simulation
Arc detector modification
300 kW RF window test
Future upgrade plan

## Slide 3
Year | RF Faults to all faults | RF MTBF
2008 | 50% | 22 Hrs
2009 | 51% | 40.5 Hrs
2010 | 64% | 41.7 Hrs
2011 | 36% | 148 Hrs
148 hrs MTBF for the year to date and still rising.
This includes all RF equipment, incl both cavities, HP amplifiers, LLRF, DAs, Cryo etc

## Slide 4
Cavity trips are the largest contributor to RF trips.
The majority of cavity trips are fast vacuum trips.
Breakdown of RF Trips

## Slide 5
Workshop on the Operation of CESR Design 500 MHz Cavities
Dates: 26-27 January 2011
SSRF
http://www.diamond.ac.uk/Home/Events/Past_events/cesr.html
Notes: 10 Attendees from other labs and manufacture
<number>

## Slide 6
Fast Cavity Vacuum Trip
Cavity field collapses within a few micro seconds.
There is always a spike on e- pickup before the trip.
Vacuum spikes observed on every gauge around the tripped cavity.
Gas can travel to and beyond the other cavity.
Gas can’t travel through the cold waveguide bend of the other cavity.
Gauge 01
Gauge 02
Gauge 03
Gauge 04
Gauge 22
Gauge 07
Gauge 08
Gauge 09
Gauge 10
No vacuum spike
Vertical scale
is auto scaling.
25µs

## Slide 7
X-ray dose rate goes up exponentially after 1.5MV.
  This indicates the existence of field emitters.
CW trip test carried out every week
  Trip at same voltage
  Conditions out quickly
  Reoccurs following 1 week of operation with beam
Gas absorption
Enhances field emission
Increases the secondary electron emission yield
Increases the risk of MP even for conditioned surfaces.
PMT
Probe
Forward power
X-ray spike during a trip
Mechanism of Fast Vacuum Trips

## Slide 8
Solutions:
Partial warm up to 35 K to release H2, annual complete warm-up
TSP pumps fired every week to increase hydrogen pumping capacity
Regular conditioning.
Vacuum system along RF Straight
Cavity 1
Cavity 3
Gauge
01
Gauge
02
Gauge
03
Gauge
04
Gauge
22
Gauge
07
Gauge
08
Gauge
09
Gauge
10
RBT
RBT
P. O. Box
FBT
FBT
P. O. Box
Partial Warm-up
Notes: Add comment regarding hydrogen being the main problem.
Increase axis title size
Explain spikes
<number>

## Slide 9
Voltage
X-ray
Spike at 1.85 MV
X-ray ~ 35 mSv/Hr
Voltage
X-ray
Steady at 1.85 MV
X-ray ~ 7 mSv/Hr
After Conditioning
X-ray Dose Rate before and after Pulse Conditioning

## Slide 10
Test of Pulse Blanker with 250mA Stored Beam
Beam oscillated but was not lost
10µs/div
20µs

## Slide 11
1 ns to 100 ns Pulse Stretcher
Adjustable Blanking
TTL Buffer
Coaxial Line driver
PXI DAQ
4x Gated Drive Amplifier Input
Block Reflected Power Protection
Prototype of Test Blanker
Waveguide
e- pickup

## Slide 12
Progress of Cavity 2 Repair
Cavity 2 was back on 04/07/2011.
We managed to install it in our RFTF in 1 week.
We started cooling down on 11/07/2011.
A big vacuum leak from helium can to cavity developed around 16 K.
The leak was later identified to be caused by feed-throughs for the probe and
waveguide e- pickup.
Glass window to view the inside of the cavity
Vacuum
CLTS sensor readings
16K
300K

## Slide 13
Cavity shape ~ Elliptic
 one point multipactor is not supported
Still there is significant chance of two point MP in the equator region in main cavity
Other MP suspected zones are – the RBT, the FBT, the coupling tongue, the RF window and the reduced height waveguide
RBT Thermal Transition
Cavity
Cu plated waveguide
FBT Thermal Transition
RBT
FBT
Nb Waveguide
Waveguide WR1800
RF Window
Coupling Tongue
Multipactor Simulations of the Diamond Cavity

## Slide 14: Modelling Multipacting in the flutes
Strong longitudinal field in the axial region
Radial near the beam tube wall and inside the flutes
Electrons in axial region can gain enough energy and shoot out of the cavity
Electrons near the beam tube wall and inside the flutes get accelerated towards the beam tube and flute wall
A transverse section through FBT at z = -162.5 mm showing electric field of the TM010 mode

## Slide 15: Multipactor in the Flutes
Multipacting in the flute for cavity voltage varied from 1.44 to 1.57 MV
V = 1.57 MV; Maximum electron energy clamped at 200 eV
Multipactor can only occur in the shoulder and is observed for high SEY material indicating a weak barrier

## Slide 16: Two Point Multipacting near the Equator
1.3 MV
1.4 MV
1.2 MV
1.5 MV
1.6 MV
1.1 MV
1.7 MV
1.8 MV
2.0 MV
2.2 MV
Number of electrons vs time (ns) for cavity voltages 1 to 2.2 MV
A well formed multipacting bunch at 1.3 MV
At low and very high voltages the rate of MP growth is lower

## Slide 17: Multipacting in the Coupling Waveguide
Cavities operated in almost matched condition
TW fields in the waveguide
Cavity SW field penetrates
into the coupling waveguide
Two point higher order MP
can occur in the waveguide
One point MP is also possible
just under the coupling tongue
The secondaries are swept away
by the TW magnetic field spreading
the MP longitudinally along the waveguide
(a) TW
(b) SW
Electric field (Abs) for input power = 0.5 W

## Slide 18: Multipacting in the coupling waveguide
Number of electrons vs time (ns) for power varied from 50 to 300 kW

## Slide 19: One Point ‘Push-Pull’ MP in the waveguide under TW conditions
28.5
29.0
29.5
30.0
30.5
31.0
Electron position plot through one RF cycle

## Slide 20: ARC Detector modifications
Suspect false trips from waveguide circulator and load arc detectors.
Examine circuit:
Very large front-end gain of 10,000,000 with no local shielding or PCB ground plane.
No filtering of inputs and outputs.
Front end op amp output shows ringing.
Tests to determine susceptibility to false trips:
Inject interference on mains cable – some small susceptibility
Interference via loop antenna onto PCB – front end op amp very susceptible
Operate mobile phone nearby – circuit fires
Sensitivity to sparks produced by gas igniter and ‘spark box’, circuit appears overly sensitive
False triggers from mobile interference
Arc detector output
Output from front end

## Slide 21: Spark box test
ARC detector
Optical fibre
Adjustable Spark gap
In-house design and build
Output voltage can be changed by adjusting the spark gap.
The total energy in the spark is determined by the capacitor bank.
Output of original arc detector saturates even with very weak sparks. A gain of 10k is enough to detect sparks of 0.1 Joule energy.

## Slide 22: Spark box results
Arc Detector Output Voltage against Spark Energy
The output was measured at different energies and gap voltages.
The gain for the desired trigger level (1 or 2 joules) is calculated using curve fitting.
The gain was reduced from 10M to 200k.
This cures susceptibility to interference.
No false trips since June ’11.

## Slide 23
300 kW RF Window Test for RI
This is to give support to industry.
System 3 was tuned up to produce 300 kW.
The 2 RF windows for PLS-II were tested to 300 kW TW and 150 kW SW.
Windows under test
RF Load

## Slide 24: Future Upgrade Plan
Installation of a 3rd cavity and increase the beam current to 500 mA.
Mezzanine floor will be installed in RF hall to provide more storage space.
New water load design is on going.
New IOT water cooling upgrade is under design.
Phase measurement and FPGA development has been progressing well.

## Slide 25
On behalf of the RF Group
Morten Jensen
Pengda Gu
Matt Maddock
Peter Marten
Shivaji Pande
Simon Rains
Adam Rankin
David Spink
Alun Watkins
Thank you for your attention!