## Slide 1: The Magnetic Measurement System Based on Coordinate Measuring Machine
CHEN, Fusan
28th June 2019

## Slide 2: The HEPS Magnet
High Energy Photon Source (HEPS)
  4th generation light source with small beam emittance.
  High gradient quadrupole/sextupole with small aperture.
  The quantity of quad. and sext. is very large.
Measurement of the high gradient magnets
All key technologies are developed and verified during the HEPS Test facility (HEPS-TF) stage

## Slide 3: The CMM based Measurement System
CMM + Rotating coil system + Stretched wire system
A multi-purpose system for magnet measurement
  Mechanical size and tolerance measurement
  Magnetic harmonics measurement
  Integral field/gradient measurement
  Magnetic field center measurement
Rotating coil system
Stretched wire system

## Slide 4: The CMM based Measurement System
System setup
CMM
Servo Motor
Angle Encoder
Tensiometer
Tension Adjustor
Motorize linear stage
High gradient quadrupole type I
Longitudinal guide
Motorize linear stage

## Slide 5: Main Component Specifications
CMM | 
Moving range [mm] | 2500×1200×800
Accuracy [μm] | 3.3+3.3L/1000
Motorize linear stage | 
Moving range [mm] | 100×100
Positioning accuracy [μm] | ±2.5
Repeatability [μm] | ±0.5
Minimum increment [μm] | 0.6
FDI2056 | 
Channel number | 3
Sample rate [kS/s] | 500
Resolution [Vs] | 10-14
Gain accuracy [ppm] | 10
Integrator drift [Fs/min] | 10-5
Power supply | 
Output voltage [V] | 60
Output current [A] | 300
Current stability [ppm] | 20

## Slide 6: Main Component Specifications
Rotating coil specifications
Skeleton material
  Diameter: 23.6mm
  Coil sag
    G10: 0.393mm
    ZrO2: 0.145mm
    Al2O3: 0.051mm
R1 [mm] | 11 | N1 | 120
R2 [mm] | 7.92 | N2 | 180
R3 [mm] | -7.84 | Coil length [mm] | 1100
R4 [mm] | -4.4 | Bucking ratio | 88

## Slide 7: Measuring Process
Create coordinate system
Measure the mechanical
Align the magnetic measurement devices
Measure the magnetic field

## Slide 8: The Coordinate System

## Slide 9: Measure the Mechanical Dimension

## Slide 10: Align the Rotating Coil

## Slide 11: Align the Stretched Wire

## Slide 12: Measurement Data

## Slide 13: Measurement Data

## Slide 14: The Vertical Magnetic Center

## Slide 15: Conclusion